## Slide 1: Research plan for graduation: Design and experimental validation of hard X-ray ionization profile monitor for Korea-4GSR
Aug. 7. 2024
Woojin Song

## Slide 2: Contents
Previous Research
Ionization Profile Monitor
Research Plan
Summary

## Slide 3: Online bunch length monitor using a fast photodiode
G. Hahn, W. Song et al,. IEEE Transactions on Instrumentation Measurement 73 (2024): 3334346
Streak Camera
High temporal resolution (< RMS 1 ps)
Slow acquisition rate
Incompatible for turn-by-turn measurements
Bunch Length
Fill-Pattern
BPM pick-up / FCT
Fill-pattern monitoring
Limited bandwidth
Moderate accuracy
Non-linear behavior causes intolerable error
Fast Photodiode
Intermediate resolution (~2 ps)
Bunch length and fill-pattern real time measurement
Bunch Length + Fill-Pattern
2 ns
470 buckets
Bunch Train,
330 bunches total
Camshaft Bunch
20 ps RMS bunch length
Sync. Radiation
MSM Photodiode
+ Bias Tee
Amp
Cable
Previous Research

## Slide 4: Online bunch length monitor using a fast photodiode
G. Hahn, W. Song et al,. IEEE Transactions on Instrumentation Measurement 73 (2024): 3334346
Experimentally demonstrated a compact monitor that enables filling pattern and bunch length monitoring using VL at the PLS-II.
Numerical compensation method to restore the original signal was shown.
 Achieved better resolution (bunch length ~ 1.2 ps, current ~ 9 μA) then conventional method
Previous Research

## Slide 5: Cherenkov diffraction radiation (ChDR) based beam diagnostics
ChDR
Charged particle
Radiator
W. Song et al,. Proc. IPAC’23
Cherenkov Diffraction Radiation (ChDR), which is emitted when charged particles pass around dielectric materials, has recently been presented as non-invasive beam diagnostics
1. Electron beam profile monitor
Non-invasive  Online measurement (OTR and YAG screens, even in storage rings).
Directional  Intuitive profile measurement, simple structure
Profile measurement experiments planned at low energy electron linac (e-LABs, later this year).
e-
ChDR
E = 100 MeV, λ = 600 nm, h = 1 mm
Previous Research

## Slide 6: Cherenkov diffraction radiation (ChDR) based beam diagnostics
2. Sub-picosecond bunch length measurement
Broad spectrum, High intensity (↔ Coherent Transition Radiation)  Non-invasive bunch length measurement using coherent ChDR spectrum
An experiment to measure electron length after acceleration in a laser wakefield accelerator is planned. (LWFA experiment using IBS 150 TW Laser)
What we meas.
(Coh. ChDR)
What we already know
(Incoh. or Single particle ChDR)
What we want know
(Bunch length)
Spectrometer
Previous Research

## Slide 7: Photon Beam Diagnostics (1/3)
In the beamline, the photon beam is delivered through various optics to the experimental end station.
The arrival position of the photon beam changes due to variations in the source position, drift (thermal, mechanical) of the optics, and vibration.
Therefore, a photon beam diagnostic is used to provide beam position and profile information for alignment to the beamline users.
4GSR Technical design report – BioPharma-BioSAXS beamline layout
HHLM: High heat load mirror, DCM: Double crystal monochromator, V/HFM: Vertical/Horizontal focusing mirror
ID beamline main component layout
Ionization Profile Monitor
Previous Research

## Slide 8: Photon Beam Diagnostics (2/3)
Photon beam diagnostic target and method
Position monitoring
  - Non-destructive : Gas type (MCP-screen, Wedge blades)
  - Semi-destructive : 4-blade type, Quadrant type
Profile monitoring
  - Non-destructive : Gas type (MCP-screen)
  - Destructive : YAG:Ce screen
The position and profile of the photon beam can be
measured non-destructively using a gas-filled chamber.
Gas
Blades
Quadrant
S.R. Marques, EPAC08
H. Aoyagi, PRAB
J. Miebner, NIMA
P. Ilinski
MCP-screen
(Optical readout)
Wedge blades
(Electrical readout)
Ionization Profile Monitor
Previous Research

## Slide 9: Photon Beam Diagnostics(3/3)
Photon beam contamination
  At the ID beamline, undulator beam is contaminated with up/down-stream bending magnet radiation, correctors, and sextuple and quadrupole with offset trajectory.*
  This contamination can cause errors in the position measurement of blade-type PBPM that signal the edge of the ID beam.
Bending
Magnet
Insertion Device
ID Beamline
BM Beamline
ID beam
BM beam
BM beam
Blade
*G. Decker, Proceedings of PAC, 1999
Ionization Profile Monitor
Previous Research

## Slide 10: Why the Gas-type Monitor?
Ionization Profile Monitor (IPM) : Visualizes the tack of the photon beam.
  Direct measurement of the beam.
  Remove position measurement errors caused by contamination.
Non-destructive profile measurement by ionizing a small fraction of the beam without blocking the photon beam.
ID beam
BM beam
BM beam
J. Kim, ID 3C SAXS Mov. Mask test, 2015
From banding magnet?
Ionization Profile Monitor
Previous Research

## Slide 11: Beam Induced Fluorescence Monitor
P. Forck, Proceedings of HB2021
Beam Induced Fluorescence monitor (BIF) - Measure the profile of an ion beam in a synchrotron.
Direct measurement of fluorescence emitted from gas molecules due to beam-gas interaction.  Long integration time required due to weak signal
Ionization Profile Monitor
Previous Research

## Slide 12: Ionization Profile Monitor (1/2)
1. Photon beam ionizes gas with a photoelectric effect.  Ion + Electron pair generation
2. Electrode transports ions to the MCP.
3. The amplified signal produces an image on a phosphor screen.
Reason why not to use electrons*
  Pros: ~8 times higher signal strength
  Cons:
  - Secondary electrons from MCP surface  Background noise
  - Higher initial velocity than ions  Increased PSF size
+
-
+
-
Photo-electron
Photo-ion
X-ray
Gas
E field
Readout (Camera)
Repeller electrode
MCP
*J. Krider, Nuclear Instruments and Methods in Physics Research Section A 278.3 (1989): 660-663.
Screen
Ionization Profile Monitor
Previous Research

## Slide 13: Ionization Profile Monitor (2/2)
Ion hits MCP channel wall and generates secondary electrons.
Electrons are again accelerated by the potential gradient and continue to multiply until they exit the channel.
Output electrons signal amplified more than 104 times in MCP are converted into a visible light on the phosphor screen.
Micro channel plate
Phosphor screen
Readout (Camera)
+
+
+
+
+
+
+
+
+
+
Visible light
+
+
e-
Del Mar Photonics
Ionization Profile Monitor
Previous Research

## Slide 14: Photon Beam IPM (1/2)
J. Miebner et al., Proceedings of the IPAC11
P. Ilinski, Proceedings of the BIW10
Using IPM for photon beam position measurement in PETRA-III (SR), FLASH (XFEL).
Calibration method and image analysis method for photon beam position and profile measurement is not proposed.
FLASH
PETRA-III (Tested at ESRF ID beamline)
Ionization Profile Monitor
Previous Research

## Slide 15: Photon Beam IPM (2/2)
Recently IPM for FEL beam monitoring at the PAL-XFEL soft X-ray beamline has been developed.
Position and beam size monitoring by measuring 500 eV mono/pink beam shot by shot.
Reduction of beam position measurement error by using Kr gas.
 Likewise, no calibration method is proposed and profile monitoring is difficult due to non-linear electric field distribution.
H. Hyun et al., “X-ray ionization beam position monitor for PAL-XFEL soft X-ray”, 2024
Ionization Profile Monitor
Previous Research

## Slide 16: Gas Jet IPM
Most synchrotrons use IPM, including HIMAC, Tevatron, RHIC, LEAR, ISR, SPS, etc.
Gas jet injection to improve reaction rate and resolution.
In the case of HIMAC, high signal intensity reduces the exposure time from 10 ms to 100 ns.
 However, complex vacuum structures are required to generate a gas jet.
Y. Hashimoto et al., Nuclear Instruments and Methods in Physics Research Section A 527.3 (2004): 289-300.
Ionization Profile Monitor
Previous Research

## Slide 17: Summary of IPM
Ionization profile monitor
  The gas in the path of the beam is ionized by the photoelectric effect, and the resulting ions are used for photon beam monitoring.
  Very low density photo-ions are guided by electrodes and amplified in the MCP to measure beam position and profile with optical readout.
Limitations of Existing Methods
  Residual gas photon IPM requires long exposure time (~ 1 sec).
  Install mesh grid in photo-ion path to make field uniform  Mesh blocks ions, reducing signal
  Calibration method not suggested  Limited to profile monitoring
What we propose:
  1. Improve signal intensity in hard X-ray
    Inject gas into the chamber instead of using residual gas
    Increase signal gain by removing mesh grid
  2. Developed a precise profile measurement method
    Install shield plate - minimize non-linear field due to removal of mesh
    Isotropic field design for precise calibration
    Development of calibration method for profile reconstruction
Ionization Profile Monitor
Previous Research

## Slide 18: Main Goal
Development of position and profile monitor for white undulator beam at Korea-4GSR
1. Gas-filled detector
  Photoionization
  Vacuum system design
  Point spread function analysis
2. Electrode design
  Optimization of isotropic field
  1:1 mapping field (defocusing field)
3. Calibration
  Magnification calibration
  Image processing
4. Experimental validation
1. Read-out device resolution
Factors to consider for improving resolution
2. Photo-ion broadening
3. Non-linear field (field quality)
4. Data analysis algorithms
Research plan
Ionization Profile Monitor
Previous Research
Research Plan

## Slide 19: 1. Gas-Filled Detector
Photoionization cross-section in the hard X-ray region of residual gas (hydrogen dominant in UHV) is significantly small.
Use of noble gases with high atomic mass to measure high energy photon.
Kr and Xe have K-absorption edges at 14 and 34 keV, respectively, so the effect on the beamline needs to be studied experimentally.
NIST X-ray form factor, attenuation, and scattering table
Hard X-ray
Soft X-ray
Ionization Profile Monitor
Previous Research
Research Plan
1. Photoionization

## Slide 20: 1. Gas-Filled Detector
Injecting gas instead of using residual gas increases chamber vacuum
 Vacuum system with differential pumping is required to minimize change in nearby beamline vacuum degree.
V. Tzoganis, Molflow+ simulation, 2017
2. Vacuum system design
Ionization Profile Monitor
Previous Research
Research Plan

## Slide 21: 1. Gas-Filled Detector
Broadening due to space charge effect of non-relativistic ions
Transverse spread due to thermal energy
Transverse spread occurs during the drift of photo-ions due to space charge and initial ion energy.
 Point spread function analysis required to determine resolution of the system.
Space charge
Ionization Profile Monitor
Previous Research
Research Plan
3. Point spread function analysis

## Slide 22: 2. Electrode Design
Mesh : Makes electric field uniform but reduces transmission due to mesh-ion collisions.
Estimated over 25% signal gain by removing mesh (W mesh open area: 64~81%).
Voltage applied to MCP lead bar creates anisotropic field  Linearity deteriorates
Mesh
Beam
+
+
+
+
+
+
+
+
+
+
Electrode
MCP
Ionization Profile Monitor
Previous Research
Research Plan
1. Isotropic field

## Slide 23: 2. Electrode Design
Install a shield plate to minimize field effects from lead potentials
PEEK plate (Insulator)
MCP
Shield plate
Equipotential line – W/O mesh
W/O mesh
W/ shield
+3 kV
-1.6 kV
-0.45 kV
Electrode
MCP
1. Isotropic field
Ionization Profile Monitor
Previous Research
Research Plan

## Slide 24: 2. Electrode Design
MCP active area (Ф 27 mm) is larger than the range the source can move (Ф 15 mm).  Design defocusing field for full MCP usage.  Achieving isotropic field simultaneously.
Beam pipe
+
+
+
+
+
+
+
+
+
2. 1:1 mapping field
Ionization Profile Monitor
Previous Research
Research Plan

## Slide 25: 2. Electrode Design
Designing an electric field optimization upgrade from an existing XFEL-XIBPM design.
  1:1 mapping field: ions generated at the edge of the expected source arrive at the edge of the MCP.
  Linearity: particles departing from the same distance spread over the same distance.  Distribution σ < 6 um achieved even at the edge of the MCP.
 Linearity must be achieved to reconstruct the original image with the calibration for profile measurement.
Initial position,
Final position,
13.5 mm
10.6 mm
Ionization Profile Monitor
Previous Research
Research Plan
2. 1:1 mapping field

## Slide 26: 3. Calibration
Transverse space @ z=0mm
Ionization Profile Monitor
Previous Research
Research Plan
1. Position calibration: Magnification correction method
V. offset = + 5 mm
V. offset = 0 mm
V. offset = - 5 mm
Defocusing field
  Lens effect distorts 2D intensity distribution information
  Magnification depends on vertical position
Center position is also measured inaccurately
 2D-Profile measurements require a calibration method to correct magnification.

## Slide 27: 3. Calibration
Real Position
Measured Position
Expected Position (Hor. Mag. )
Exp. Position (Ver. Mag. )
Magnification correction method
2 directions (Vertical, Horizontal) monitor required.
Move beam in X, Y direction to measure calibration map (𝑀).
Calculate actual position () by substituting the measured beam center (, ).
Divide the position information of the MCP image by the magnification (M) for profile reconstruction
Hor. monitor
Ver. monitor
Ionization Profile Monitor
Previous Research
Research Plan
1. Position calibration: Magnification correction method

## Slide 28: 3. Calibration
Hor. Monitor
(offset: 3, 5 mm)
Ver. Monitor
(offset 5, -3 mm)
Initial peak position:
Calibrated peak position:
Calibration method can restore initial position and original profile with magnification correction
Ionization Profile Monitor
Previous Research
Research Plan
1. Position calibration: Magnification correction method

## Slide 29: 3. Calibration
Position measurement error of Magnification correction method is within 40 μm (MCP resolution ~ 80 μm).
Calibration error can be further reduced by using higher order terms in the equation for magnification curve ().
Need to verify calibration method by measuring calibration map experimentally.
Ionization Profile Monitor
Previous Research
Research Plan
1. Position calibration: Magnification correction method

## Slide 30: 3. Calibration
Screen-camera alignment
  Applying a correction method to find the center and beam path of a screen image.
Noise removal
  Image process algorithm for removing background noise to improve resolution
Ionization Profile Monitor
Previous Research
Research Plan
2. Image processing method

## Slide 31: 4 Experimental validation
2024 Photon IPM Development Schedule
Detail Items | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
Basic study |  |  |  |  |  |  |  |  |  |  |  | 
Main component (MCP, power supply) order |  |  |  |  |  |  |  |  |  |  |  | 
Electrode design |  |  |  |  |  |  |  |  |  |  |  | 
Vacuum chamber design |  |  |  |  |  |  |  |  |  |  |  | 
Calibration method development |  |  |  |  |  |  |  |  |  |  |  | 
Fabrication and assembly |  |  |  |  |  |  |  |  |  |  |  | 
PLS-II (6C) ID Beamline installation |  |  |  |  |  |  |  |  |  |  |  | 
Calibration and test |  |  |  |  |  |  |  |  |  |  |  | 
Ionization Profile Monitor
Previous Research
Research Plan

## Slide 32: 4 Experimental validation
Magnification map
  Measurement position changes as a actual beam position changes
  Calculate magnification by moving the IPM chamber horizontally and vertically
  Verifying the difference depending on the calibration beam size
Experimental validation
  Photon energy
  White beam, Mono-energetic beam
  Gas species and vacuum degree
  Residual gas, Krypton, Nitrogen
  Effect on the experimental beam due to K-absortion edges
Magnification vs Cal. Beam size
Ionization Profile Monitor
Previous Research
Research Plan

## Slide 33: Summary
Developing an Ionization Profile Monitor to measure the profile and position of the white undulator beam for Korea-4GSR.
Unlike blade PBPM, IPM is non-destructive and not affected by contamination.
Existing photon IPM are limited in profile measurement because there is no calibration method.
Low signal intensity is expected from residual gas, so plan to improve signal gain by injecting noble gas.
Designed a defocusing field to utilize all of the active area of the MCP.
Suggested a new calibration method and image processing method to measure accurate profiles.
The new proposed design will be experimentally validated on the PLS-II.
Ionization Profile Monitor
Previous Research
Research Plan
Summary

## Slide 34: Appendix

## Slide 35
Ionization Profile Monitor
Previous Research
Research Plan
Summary
Appendix