## Slide 1: Introduction of BPM research & Deflecting cavity verification
Geunwoo Kim
2024. 08. 14.

## Slide 2: Contents
Beam position monitor (BPM) research
  BPM signal & position definition (RAON)
  BPM offline calibration
RF deflecting cavity simulation (CST)

## Slide 3: Beam position monitor, BPM
Non-destructive diagnostics device
Enable real time beam position measurement
  +) Time of flight, phase, etc.
Beam
E
Ibeam(t)
Iimage(t)
<Equivalent circuit>

## Slide 4: Beam position monitor, BPM
① Measurement
② Signal processing
③ Estimation of beam position
Four electrodes

## Slide 5: BPM signal processing procedure
ADC
BPM offline calibration research
Broadband
Narrowband
Signal strength
definitions
BPM
Position
Calculation
definitions
Analog
Front-end
stage
Top
Right
Left
Bottom
Data selection for calibration
Position calibration
BPM signal & position definition (RAON)
Vright, Vleft
Vtop, Vbottom
Xbpm, Ybpm
Xwire = K*Xbpm + offset

## Slide 6: Filtering process
ADC
Signal strength
definitions
BPF
Position
Calculation
definitions
Broadband processing
1st freq.: 81.3 MHz
2nd freq.: 162.6 MHz
Narrowband processing
ADC
Broadband
Narrowband
Signal strength
definitions
BPM
Position
Calculation
definitions
Analog
Front-end
stage
Position
calculation
definitions
ADC
Signal strength
definitions
Full range signal
(No filtering)
Discussions
Broad vs Narrow
Signal definition
Position calculation
  Calibration factor
  Error/Accuracy
A
Noise 10%
Noise 10%
Data: CST simulation
Vright, Vleft
Vtop, Vbottom
Xbpm, Ybpm

## Slide 7: BPM pickup signal strength definitions
Defining signal strength is essential to use digitized ADC values as representative electrode signals.
How should the digitized ADC count values be defined regarding signal strength?
 Define five different types of signal strength and compare their characteristics.
Accumulation (ACCU) ()
Mean square error (MSE) ()
Root mean square error (RMSE) ()
Mean absolute error (MAE) ()
Peak ()
N: # of samples
: ith ADC count
: Averaged ADC count value
ADC
Broadband
Narrowband
Signal strength
definition
BPM
Position
Calculation
definition
Analog
Front-end
stage
A
B
C
D
Vright, Vleft
Vtop, Vbottom

## Slide 8: Characteristics of each signal strength definition
Accumulation (ACCU) ()
Mean square error (MSE) ()
Root mean square error (RMSE) ()
Mean absolute error (MAE) ()
Peak ()
For MSE, when the error is less than 1, it shows a lower error compared to other signal strengths. RMSE and MAE exhibit a linear increase in error.
|Error| ≤ 1
ADC
Broadband
Narrowband
Signal strength
definition
Position
Calculation
definition
Vright, Vleft
Vtop, Vbottom
Xbpm, Ybpm

## Slide 9: BPM position calculation definitions
Using 5 types of signal strength and 3 position calculation formulas, compare and analyze the correction coefficient K and error values.
ACCU
RMSE
MAE
MSE
PEAK
Δ/Σ
Log
Atan
Signal strength definitions
Position calculation fomulas
Difference-over-sum (DOS)
  Delta-over-sigma,
  Log,
  Arctan,
Position calibration method: linear fitting x: a certain DOS value
              K: calibration coefficient
              δ: offset
Coefficient K ∝ Signal noise
∴ Lower K is better
Vleft
B
Signal
Strength
definitions
Vright
A
Position
Calculation definitions

## Slide 10: Broadband processing
Difference-over-sum
Calibration range

## Slide 11: Broadband processing
Linear fitting (Calibration range ± 5 mm2)
Y=X graph and calibration results plot
<Residual>
Target - corrected

## Slide 12: Comparison of 15 different calibration results
ADC
Broadband
Narrowband
Signal strength
definition
Position
Calculation
definition
Broad vs Narrow: With simulation + noise data, there is no significant impact on the calibration coefficients and errors between them.
Signal strength: Unless defined as in the case of MSE, similar results are observed.
Position calculation: It has been confirmed that it primarily influences the calibration coefficients and errors.
ACCU
RMSE
MAE
MSE
PEAK
Δ/Σ
Log
Atan
LogMSE is the most suitable for position correction among the 15 methods.
Signal strength definitions
Position calculation fomulas
Position calibration
Axis data was selected
Data selection for calibration

## Slide 13: BPM signal processing procedure
ADC
BPM offline calibration research
Broadband
Narrowband
Signal strength
definitions
BPM
Position
Calculation
definitions
Analog
Front-end
stage
Top
Right
Left
Bottom
Data selection for calibration
Position calibration
BPM signal & position definition (RAON)
Vright, Vleft
Vtop, Vbottom
Xbpm, Ybpm
Xwire = K*Xbpm + offset

## Slide 14: Used BPM for offline calibration
Rectangular shape electrode (similar with stripline BPM, but not stripline BPM)
The BPM was designed as part of a plan to measure various particles.
  200 keV/u, 65 MHz
  1.5 MeV/u, 200 MHz
  10 MeV/u, 352 MHz
 In research, we performed a BPM offline calibration targeting 10 MeV/u, 352 MHz.

## Slide 15: BPM offline calibration
Does not use actual beam signals for calibration
Utilizes RF signals applied to wire to simulate the beam.
The BPM is calibrated using this simulated beam method.

## Slide 16: Why need a BPM test bench?
Use of multiple BPMs
Custom-developed BPM for their specific needs
Manufacturing tolerances
Offset Identification.
Ensuring proper calibration characteristics (calibration coefficient, etc.).
Data selection for calibration
Position calibration
Xwire = K*Xbpm + offset

## Slide 17: BPM offline calibration procedure
Difference-over-sum (DOS)
  Delta-over-sigma,
  Log,
  Arctan,
Data selection for calibration
Position calibration
Xwire = K*Xbpm + offset

## Slide 18: BPM offline calibration results
Calibration using selected DOS data on the ±(8,8) mm2 range
  1, 3, 5th, and 2D-3rd polynomial fit correction
  1D polynomial:
  2D polynomial:

## Slide 19: BPM offline calibration results
2D colormap

## Slide 20: BPM offline calibration results
3rd fit correction ≤ 100 μm accuracy
DOSequal: ~ ±(6,6) mm2
DOSaxis : ~ ±(7,7) mm2
DOSall : ~ ±(3,3) mm2
2D-3rd: ±(5,5) mm2
3rd fit correction @ ±(8,8) mm2, 100K samples
X STD ≤ 10 μm
Y STD < 35 μm
Whole range ≤ 100 μm STD

## Slide 21: Contents
Beam position monitor (BPM) research
  BPM signal & position definition (RAON)
  BPM offline calibration
RF deflecting cavity simulation (CST)

## Slide 22: RF deflecting cavity simulation (CST)
RF
deflecting
cavity
E0
Beam
E0
eVt
x’
Drift space
Temporal beam diagnostics for e-labs
Measurement of the longitudinal bunch length of the electron beam.
E-labs deflector

## Slide 23: RF deflecting cavity simulation (CST)
Eigenmode solver: frequency mismatching  solved
Frequency solver: S11, S21 curve mismatchings  still progressing
Common: change mesh properties (model mesh 24 → 40)

## Slide 24: Eigenmode solver: frequency mismatching
<Simulation>
<Reference>
 | Previous sim. | Current sim. | Reference
Frequency [GHz] | 2.86519 | 2.85601 | 2.85604
Phase [deg] | 0 | 210 | 210
Max E(Plane) [V/m] | 4.45e7 | 4.76e7 | 4.67e7
Mesh parameters | 20 | 40 | ?? (maybe auto)
∴ Mesh was not enough to measure freq. properly.

## Slide 25: frequencymode solver
<Previous sim.>
<Current sim.>
<Reference>
Change the # of data samples (1,001 → 100,001)
Need more mesh values?

## Slide 26: Next plans
Environment setting at POSTECH (CST server, PC, …)
CST simulation
  Change Mesh properties (matching results of freq & time solver)
  RF deflecting cavity theory study w/ formulas