## Slide 1: Research Progress
Nov. 01. 2024
송우진

## Slide 2
Review
  Beam induced fluorescence monitor
  Ionization profile monitor
  Gas-jet type IPM
  IPM for photon beam monitoring
PAL-XFEL XIBPM
XIBPM upgrade
  1:1 mapping field design
  Shield plate
Summary of GPBPM Development
1. Background Study
4GSR-GPBPM
  Electrode structure design
  Target 1: 1:1 mapping field
    Electrode modification
    Simulation results for several parameters (applied electrode voltage, inner diameter, position)
  Target 2: Achieving linearity
    Shield plate modification
    Repeller size, electrode outer diameter change
Calibration
  Issue
  Calibration method
  Case study
    Gaussian distribution (W/O beam pipe)
    Part of Gaussian (W/ beam pipe)
    Mixture of Gaussian distribution
    Distribution with flat-top
  Calibration error estimation
  Calibration beam size – slit (iris) size
    Calibration map based on calibration beam size
    Error analysis due to calibration beam size
2. Conceptual Design

## Slide 3
Magnification modification
  Target magnification
  Result
Determination of main parameters
Manufacturing drawing
3. Detailed Design
Expected signal strength
  Photon-gas interaction
  Spectral photon flux
  Event number estimation
Transverse spread
  Space charge effect
  Thermal energy
Systematic error
  Resolution estimation
  Experimental results
5. Resolution
Experimental plan
  Experimental goal and contents
  Preliminary preparations
  Detailed experimental procedure
Fabrication
Test
  12/16-19 PLS-II BL-6C
4. Experiment
Summary of GPBPM Development

## Slide 4
Main goal: - Evaluate the beam profile and position measurement performance of the PBPM
Contents
  Gas pressure optimization (minimum injection)
    TMP – Exp. condition, independent system / Ion pump – Beamline, connected system
  Beam size and position variation: 0.1 ~ 6 mm, Ф 12 mm (*Microscope FOV: 4.16 mm (H) @ 4x, 8.3 mm (H) @ 2x)
  (Optional) Beam energy: mono beam (10 ~ 40 keV)
  (Optional) Photon flux variation
Experimental Plan

## Slide 5
Layout of PLS-II BL-6C
Multi-Pole Wigger
Ec = 12 keV
Normal source size
: 500 um x 30um
Total power : 10kW
36m
Double Multilayer
Monochromator
Ru/C, 4.0 nm
: 10 – 20 keV
W/B4C, 2.6 nm
: 20 – 55 keV
DCM 1
Si (220)
: 30 – 50 keV
DCM 2
Si (111)
: 12 – 33 keV
Sample stage
Air bearing rotation stage
sCMOS Camera
Pco.Edge 5.5
Hamamatsu Fusion
2x, 5x, 10x, 20x, 40x
Slit
LMM

## Slide 6
Chamber manufacturing and vacuum test
DAQ (Camera) interface
Gas injection test
BL-6C white beam ray-tracing
(Optional) Confirm applicability of motorized stage  Replace with beam adjustment
(Optional) Cooperation with Survey/Alignment team  Ray-tracing (Beam based alignment)
Procedure
Completion of the chamber  Preliminary test (Gas, Pumping, Beam tuning, etc.)  White beam ray-tracing  Installing in BL-6C  PBPM and beam alignment  Experiment: 12/16 (Mon) to 12/19 (Thu)
Preliminary preparations

## Slide 7
Gas pressure optimization
  Measure signal intensity and PSF according to gas species (Xe, Res. gas) and pressure.
  Electrical control of gas pressure (MFC) is required? (Stability, automation of the measurement process).
Beam size and position measurement
  Measure performance of beam size (profile) and position monitoring.
  Verify magnification correction method (Calibration method).
  Measure point spread function.
Detailed experimental procedure
Window (Φ12 mm)
Beam size 0.1 ~ 6 mm
Position
PSF measurement

## Slide 8
3. (Optional) Beam energy: Mono energetic beam profile measurement.
4. (Optional) Photon flux: Test signal intensity changes proportional to beam intensity. → Verify the applicability of the beam intensity monitor.
Detailed experimental procedure
1st harmonic (1955 eV)
20 keV

## Slide 9
Expected signal strength
  Photon-gas interaction
  Spectral photon flux
  Event number estimation
Transverse spread
  Space charge effect
  Thermal energy
Systematic error
  Resolution estimation
  Experimental results
Resolution estimation

## Slide 10
Up to approximately 100 keV photon energy, the photoelectric effect is predominantly governing.
Beyond 100 keV, Compton scattering and pair production become dominant, but these are outside the region of interest.
Photon-gas interaction
τ: Photoelectric effect
: Rayleigh scattering : Compton scattering : Photonuclear reaction : Pair production

## Slide 11
Photoionization cross-section
Single photoionization cross-section
Charge state spectrum of Xe in collision with 4.5 and 5.6 keV photons

## Slide 12
Spectral photon flux - Spectra
4GSR IVU20 (8 mm x 8 mm @ 40m, W/O diffraction)

## Slide 13
Spectral photon flux - Spectra
4GSR IVU20 (8 mm x 8 mm @ 40m, W/O diffraction)
Total flux (4 pi) including off-axis radiation
On-axis flux density

## Slide 14
Event number estimation
Number of interaction
= Cross section *
Spectral photon flux *
Gas density (1e-6 Torr)
 | PLS-II IVU20 | 4GSR IVU20
Total event (0.1~50 keV) | 3.00e+9 | 7.91e+9
@ 10 keV (0.1% B.W.) | 1.85e+5 | 2.50e+6
@ 30 keV (0.1% B.W) | 5.09e+2 | 2.59e+4

## Slide 15
Space charge effect

## Slide 16
The space charge model for IPM has been calculated experimentally, but the beam size analyzed in this paper provides a more accurate calculation.
Space charge effect

## Slide 17
Measurement beam size is calculated by initial beam size ,
1. ,
2. , where
Space charge effect
Beam
Detector
Parameters | Description | Scale (Our case)
 | RMS bunch length | 20 ps
 | Time between bunches | 2 ns
 | Time for the secondaries to get extracted out of the beam | 78 ns
 | Expansion time due to space charge | 20 us
 | Time until the particle reaches the detector | 1 us
Parameters | Description | Scale (Our case)
 | RMS bunch length | 20 ps
 | Time between bunches | 2 ns
 | Time for the secondaries to get extracted out of the beam | 78 ns
 | Expansion time due to space charge | 20 us
 | Time until the particle reaches the detector | 1 us

## Slide 18
(Assuming the is constant)
Thermal velocity
 Independent of gas type , RMS , RMS
Transverse spread due to thermionic energy
Parameters | Description
 | Transverse displacement
 | Vertical drift length (40 mm)
 | Time of flight
 | Atomic mass (Xe: 131 amu)
 | Charge number (+1, +10)
Parameters | Description
 | Transverse displacement
 | Vertical drift length (40 mm)
 | Time of flight
 | Atomic mass (Xe: 131 amu)
 | Charge number (+1, +10)
Defocusing field

## Slide 19
Calculate the transverse spread numerically using a particle tracking code. (Space charge effect is not considered)
Transverse spread simulation

## Slide 20
Transverse spread due to thermal energy
- 155 um (for +1), 49 um (for +10)
The tracking results show a 15% greater broadening compared to analytic calculations.
In the tracking simulation, the actual (defocusing) field is implemented.
 Particles with transverse momentum are accelerated in the transverse direction the moment they deviate from the center.
Transverse spread simulation

## Slide 21
An experiment plan for the ionization profile monitor has been established.
While waiting for its fabrication, aim to calculate the resolution of the IPM.
Resolution of device
Summary
Factor | Source | Quantity ()
Field quality | Non-linearity | < 6 um
 | Defocusing field | (Magnification)
Transverse spread | Thermal energy | ~ 60 um (?)
 | Space charge effect | 2.1 um (@ 3 mm)
Readout device resolution | MCP | ~ 80 um
 | Phosphor screen | ~ 80 um
Data processing | Calibration method | < 40 um
 | Fitting error (Dep. on camera) | TBD
Factor | Source | 
Field quality | Non-linearity | < 6 um
 | Defocusing field | 
Transverse spread | Thermal energy | ~ 60 um (?)
 | Space charge effect | 2.1 um (@ 3 mm)
Readout device resolution | MCP | ~ 80 um
 | Phosphor screen | ~ 80 um
Data processing | Calibration method | < 40 um
 | Fitting error (Dep. on camera) | TBD

## Slide 23
…
monitor
1 mm
1.1 mm
1.05
1.05
Horizontal center
Horizontal monitor
monitor
Vertical monitor
Vertical center
Alignment using ray-tracing